## Slide 1: Cryogenic Interferometric Inertial Sensors for next-generation Gravitational-Wave Detectors
Morgane Zeoli, T. Giordano, B. Thibaut, M. Teloi, L. Jacques,
G. Bruno, J. van Heijningen, C. Collette
Belgian-Dutch Gravitational-Wave Meeting 2026 – Louvain-la-Neuve

## Slide 2: Goal and requirements
E-TEST is a prototype suspension, combining passive and active isolation techniques, of a 100 kg silicon mirror cooled down radiatively to 25 K in a suspended cryostat
Develop and validate cryogenic compatible inertial sensors with a sensitivity of m/√Hz from 2-3 Hz at 20K
Typical ground motion
Target residual motion at E-TEST mirror level
Monitor residual motion at the last isolation stage, in the cryostat
Commercial sensor (not vacuum nor cryo compatible)
PML legacy (not cryo compatible)
VINS (Binlei Ding): PhD thesis
EVINS (Anthony Amorosi): G2502591
Objective resolution for the cryogenic sensor

## Slide 3: High-resolution cryogenic inertial sensor suite combining passive and active architectures
Interferometric readout
Vertical cryogenic inertial sensor
Horizontal cryogenic inertial sensor
Passive sensing
Active sensing
Homodyne quadrature interferometric readout
Amorosi 2025
Leaf-spring suspension (BeCu)
Double-clamped blade hinge (BeCu)
Inertial mass
Corner cube
Frame
Monolithically suspended Ti6Al4V mass
Voice-coil actuator
Current injector
Not cryogenic
Cryo-V
Cryo-H

## Slide 4: Homodyne quadrature interferometric readout sub pm/√Hz resolution from 0.1Hz at 293K
Null measurement: the corner cube is clamped to the interferometer to cancel relative motion between the measuring and reference arm
Limited by dark current noise
Shot noise: ~50 fm/√Hz
Homodyne quadrature interferometer
Remove the remaining vibration by removing the coherent signals
Fixed!

## Slide 5: Homodyne quadrature interferometric readout sub pm/√Hz resolution from 0.4Hz at 17K
Alignment quality: less than 10% of relative fringe visibility loss from 293K to 17K
Can be used as a standalone cryogenic sub-pm/√Hz sensitivity relative displacement sensor
Thermal-induced drift
Electronics limited
Shot noise: ~50 fm/√Hz

## Slide 6: Cryo-V gravity-compensation spring: Elinvar alloys decrease the sensor sensitivity to temperature
Thermal treatment to get an almost-zero TEC
Drift(T)
increase
Have a suspension with a low thermoelastic coefficient (TEC) and low coefficient of thermal expansion (CTE)
Decrease measurement bandwidth
Nispan-C limits the sensor sensitivity to thermal fluctuations
Combine materials with opposite TEC

## Slide 7: Cryo-V operation at 5.17 K compared to commercial sensors
T =
= Coherent from 0.1 to 30Hz
= Coherent from 12 to 50Hz
Vacuum tank
4K Cold plate
Guralp
Accelerometer
Cryo-Vs
T =
Cryo-V operation assessed at 5K from 0.1 to 50Hz

## Slide 8: Cryo-H actuation chain adaptation to cryogenic temperature
VCA with copper coil and neodymium magnet
Low-noise voltage buffer made with MELF resistors

## Slide 9: Cryo-H operation at 27.17 K compared to commercial sensors
Coherent from 20 mHz to 150Hz
T =
Guralp
Cryo-H
Test bench dynamics
Cryo-H operation assessed at 27K from 20mHz to 150Hz
Homemade low-vibration cryogenic testbench

## Slide 10: Horizontal cryogenic inertial sensor resolution at 75 K reaches sub-pm level from 5 Hz
Shot noise < 10 fm/√Hz
Actuator noise
-
Huddle test
= coherent subtraction
Capacitive and cross coupling
(alignment- and input-power-dependent)
Mitigated by decreasing the actuator strength

## Slide 11: Conclusion and ongoing activities
We have redesigned and adapted state-of-the-art sensors to cryogenic conditions
Checked the cryogenic compatibility of sub-components
The homodyne quadrature interferometer can be used as a standalone cryogenic sub-pm/√Hz sensitivity relative displacement sensor
The vertical cryogenic inertial sensor is operating at 5K from 0.1 to 50Hz
The horizontal cryogenic inertial sensor is operating at 27K from 30 mHz to 150Hz with a sub-pm resolution from 5 Hz
Integration in large facilities (E-TEST)

## Slide 12: Backup slides

## Slide 13: Low-vibration cryogenic testbench
Active platform
Geophones (custom Theta from GaiaCode)
Suspension and actuator
CS210-GMX-20B
cryocooler
Cryocooler support
Rubber bellow
gHe-filled heat exchanger
Vacuum chamber
Cryocooler 45K stage
Intermediate thermal shield
Soft OFC thermal braids
Cryocooler 4K stage
Cold thermal shield
10 nm RMS in [0.01-10] Hz
pm/√Hz level from 3 Hz
15-20 K

## Slide 14: Interferometric readout cryogenic compatibility: Critical optical elements selection for 1550nm
Kulbusch, Zeoli et al (2024)
Relative decrease in voltage <2.2%
Collimators
Homodyne quadrature interferometer

## Slide 15: Interferometric readout cryogenic compatibility: Critical optical elements selection for 1550nm
Kulbusch, Zeoli et al (2024)
Photodiodes
Homodyne quadrature interferometer
InGaAs ++ without front glass
InGaAs +
InGaAs
InGaAs ++
Bajpai (2022)

## Slide 16: Low-vibration cryo testbench
19K
1 order of magnitude at 0.1 Hz

## Slide 17: Active isolation

## Slide 18: Vertical cryogenic inertial sensor
B. Ding, G. Zhao, J. Watchi, A. Sider and C. Collette, “An interferometric inertial sensor for low-frequency sensing,” Sensors and Actuators: A. Physical 335, 2021.
Common-mode noise rejection
Online-ellipse fitting and normalization
Fringe demodulation

## Slide 19: Interferometric readout – Alignment & compatibility
Alignment quality figure of merit:
Fringe visibility (FV) of the raw fringes measured by each photodiode
Cryo-V
Temperature sensors
Blocking rod
Graphite layer
Thermal link
Interferometric readout
The interferometer alignment is maintained down to 20K
Fitted and normalized ellipse
  <10% relative fringe visibility loss!